# Table S1: Scan parameters

**High resolution reference scan**

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| **Scan Series** | **Tube-voltage Combination (kV)** | **Rotation Time (s)** | **Pitch** | **Tube Current (mAs)** | **CTDIVol** | **Collimation** |
| 0\_0a) | 120 | 1,0 | 0,2 | 685 | 40,11 | 64 x 0,6 |

**Experiment 1 and 2: Calibration, Dose and Tube Voltage**

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| **Scan Series** | **Tube-voltage Combination (kV)** | **Rotation Time (s)** | **Pitch** | **Tube Current (mAs)** | **CTDIVol** | **Collimation** |
| 1\_1 | 70 / Sn150 | 0,5 | 0,7 | 44 | 1,02 | 128 x 0,6 |
| 1\_2 | 70 / Sn150 | 0,5 | 0,7 | 131 | 3,02 | 128 x 0,6 |
| 1\_3 | 70 / Sn150 | 0,5 | 0,7 | 218 | 5,01 | 128 x 0,6 |
| 1\_4 | 70 / Sn150 | 0,5 | 0,7 | 473 | 10,00 | 128 x 0,6 |
| 1\_5 | 70 / Sn150 | 0,5 | 0,7 | 875 | 20,00 | 128 x 0,6 |
| 1\_6 d) | 70 / Sn150 | 1,0 | 0,3 | 3500 | 79,93 | 128 x 0,6 |

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| **Scan Series** | **Tube-voltage Combination (kV)** | **Rotation Time (s)** | **Pitch** | **Tube Current (mAs)** | **CTDIVol** | **Collimation** |
| 2\_1 | 80 / Sn150 | 0,5 | 0,7 | 24 | 1,01 | 128 x 0,6 |
| 2\_2 | 80 / Sn150 | 0,5 | 0,7 | 73 | 3,04 | 128 x 0,6 |
| 2\_3 | 80 / Sn150 | 0,5 | 0,7 | 121 | 5,02 | 128 x 0,6 |
| 2\_4 | 80 / Sn150 | 0,5 | 0,7 | 243 | 10,03 | 128 x 0,6 |
| 2\_5 | 80 / Sn150 | 0,5 | 0,7 | 486 | 20,00 | 128 x 0,6 |
| 2\_6 d) | 80 / Sn150 | 1,0 | 0,3 | 1945 | 80,16 | 128 x 0,6 |

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| **Scan Series** | **Tube-voltage Combination (kV)** | **Rotation Time (s)** | **Pitch** | **Tube Current (mAs)** | **CTDIVol** | **Collimation** |
| 3\_1 | 90 / Sn150 | 0,5 | 0,7 | 16 | 1,02 | 128 x 0,6 |
| 3\_2 | 90 / Sn150 | 0,5 | 0,7 | 48 | 3,00 | 128 x 0,6 |
| 3\_3 | 90 / Sn150 | 0,5 | 0,7 | 81 | 5,04 | 128 x 0,6 |
| 3\_4 | 90 / Sn150 | 0,5 | 0,7 | 162 | 10,03 | 128 x 0,6 |
| 3\_5 | 90 / Sn150 | 0,5 | 0,7 | 323 | 20,00 | 128 x 0,6 |
| 3\_6 d) | 90 / Sn150 | 1,0 | 0,3 | 1295 | 80,12 | 128 x 0,6 |

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| **Scan Series** | **Tube-voltage Combination (kV)** | **Rotation Time (s)** | **Pitch** | **Tube Current (mAs)** | **CTDIVol** | **Collimation** |
| 4\_1b) | 100 / Sn150 | 0,5 | 0,7 | 24 | 1,50 | 128 x 0,6 |
| 4\_2 | 100 / Sn150 | 0,5 | 0,7 | 48 | 3,01 | 128 x 0,6 |
| 4\_3 | 100 / Sn150 | 0,5 | 0,7 | 80 | 4,97 | 128 x 0,6 |
| 4\_4 | 100 / Sn150 | 0,5 | 0,7 | 161 | 10,05 | 128 x 0,6 |
| 4\_5 | 100 / Sn150 | 0,5 | 0,7 | 321 | 20,00 | 128 x 0,6 |
| 4\_6 d) | 100 / Sn150 | 1,0 | 0,3 | 1290 | 80,12 | 128 x 0,6 |

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| **Scan Series** | **Tube-voltage Combination (kV)** | **Rotation Time (s)** | **Pitch** | **Tube Current (mAs)** | **CTDIVol** | **Collimation** |
| 5\_-b) | - | - | - | - | - | - |
| 5\_2 | 80 / 140 | 0,5 | 0,7 | 66 | 3,07 | 128 x 0,6 |
| 5\_3 | 80 / 140 | 0,5 | 0,7 | 110 | 5,00 | 128 x 0,6 |
| 5\_4 | 80 / 140 | 0,5 | 0,7 | 223 | 10,00 | 128 x 0,6 |
| 5\_5 | 80 / 140 | 0,5 | 0,7 | 445 | 20,03 | 128 x 0,6 |
| 5\_6 d) | 80 / 140 | 1,0 | 0,3 | 1780 | 80,00 | 128 x 0,6 |

**Experiment 3: RotationTine**

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| **Scan Series** | **Tube-voltage Combination (kV)** | **Rotation Time (s)** | **Pitch** | **Tube Current (mAs)** | **CTDIVol** | **Collimation** |
| 6\_1 | 80 / Sn150 | 0,25 | 0,7 | 121 | 5,00 | 128 x 0,6 |
| 6\_2 | 80 / Sn150 | 0,28 | 0,7 | 121 | 5,00 | 128 x 0,6 |
| 6\_3 | 80 / Sn150 | 0,5 | 0,7 | 121 | 5,02 | 128 x 0,6 |
| 6\_4 | 80 / Sn150 | 1,0 | 0,7 | 121 | 5,02 | 128 x 0,6 |

**Experiment 4: Pitch**

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| **Scan Series** | **Tube-voltage Combination (kV)** | **Rotation Time (sec)** | **Pitch** | **Tube Current (mAs)** | **CTDIVol** | **Collimation** |
| 7\_1 | 80 / Sn150 | 0,5 | 0,3 | 121 | 5,11 | 128 x 0,6 |
| 7\_2 | 80 / Sn150 | 0,5 | 0,5 | 121 | 5,03 | 128 x 0,6 |
| 7\_3 | 80 / Sn150 | 0,5 | 0,8 | 121 | 5,00 | 128 x 0,6 |
| 7\_4 | 80 / Sn150 | 0,5 | 1,2 | 121 | 5,02 | 128 x 0,6 |

**Experiment 5: Reconstruction Kernel**

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| **Scan Series** | **Tube-voltage Combination (kV)** | **Rotation Time (s)** | **Pitch** | **Tube Current (mAs)** | **CTDIvol** | **Collimation** |
| 8\_1c) | 80 / Sn150 | 0,5 | 0,7 | 121 | 5,02 | 1. 0,6
 |

Scanned on Siemens Somatom Force on the 15.08.2018 at Herlev Hospital. Scans were reconstructed using kernel Qr40 with a slice thickness of 0.75 mm and a spacing of 0.5 mm a) Scanned using ultra high resolution scan mode b) at a Tube-voltage combination of 100 kV / Sn150 kV and 80 kV / 140 kV the scanner could not be reduced in CTDIvol to 1 mGy c) Reconstructed using kernels Qr32 - Qr36 - Qr40 - Qr44 - Qr49 - Qr54 - Qr59 - Qr69 d) scans used for calibration of the three material decomposition;

*CTDIvol* volume computed tomography dose index; *Sn* Indicates use of 0.6mm tin filter;

# Table S2: All experimental results

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| **S2a) Dose & Tube-voltage (image noise)** |  |  |  |  |  |
|  | ***Tube voltage combination*** |  |  |  |  |  |
| ***CTDIVol (mGy)*** | ***70 kV / Sn150 kV*** | ***80kV / Sn150 kV*** | ***90 kV / Sn150 kV*** | ***100 kV / Sn150 kV*** | ***80 kV / 140 kV*** |  |  |  |  |  |
| ***1*** | 69.6 | 74.5 | 85.9 | 85.9 a) | - b)  |  |  |  |  |  |
| ***3*** | 41.6 | 45.1 | 50.7 | 61.8 | 78.3 |  |  |  |  |  |
| ***5*** | 33.6 | 35.1 | 40.7 | 47.8 | 61.1 |  |  |  |  |  |
| ***10*** | 25.8 | 27.2 | 31.0 | 35.8 | 45.6 |  |  |  |  |  |
| ***20*** | 20.9 | 21.9 | 24.6 | 27.5 | 37.1 |  |  |  |  |  |
| ***80*** | 14.6 | 14.7 | 15.9 | 16.6 | 19.9 |  |  |  |  |  |
| *Tube voltage combination* | *rs* -0.98 | *p < 0.001* |  |  |  |  |  |  |
| *Dose (CTDIvol)* | rs -0.96 | *p < 0.001* |  |  |  |  |  |  |
|  |  |  |  |  |  |
| **S2b) Dose & Tube-voltage (mean error)** |  |  |  |  |  |
|  | ***Tube voltage combination*** |  |  |  |  |  |
| ***CTDIVol (mGy)*** | ***70kV/ Sn150kV*** | ***80kV/ Sn150kV*** | ***90kV/ Sn150kV*** | ***100kV/ Sn150kV*** | ***80kV/ 140kV*** |  |  |  |  |  |
| ***1*** | 10.0 | 14.0 | 8.5 | 9.0 a) | - b) |  |  |  |  |  |
| ***3*** | 10.6 | 14.1 | 11.1 | 11.5 | 11.9 |  |  |  |  |  |
| ***5*** | 10.0 | 14.2 | 11.7 | 10.5 | 11.5 |  |  |  |  |  |
| ***10*** | 11.4 | 13.9 | 12.8 | 8.2 | 14.7 |  |  |  |  |  |
| ***20*** | 9.5 | 14.7 | 13.9 | 9.8 | 13.5 |  |  |  |  |  |
| ***80*** | 10.5 | 12.7 | 11.7 | 10.9 | 14.0 |  |  |  |  |  |
| Dose (CTDIvol) | *rs 0.17* | p = 0.365 |  |  |  |  |  |  |
|  |  |  |  |  |  |
|  |
| **S2c) Rotation Time** |
|  | ***Rotation Time*** |  |  |  |  |  |  |
|  | *0.25 s* | *0.28 s* | *0.5 s* | *1.0 s* | ***rs*** | ***p*** |  |  |  |  |
| ***Mean Error*** | 14.5 | 14.5 | 9.5 | 14.1 | -0.80 | 0.333 |  |  |  |  |
| ***Image Noise*** | 36.8 | 36.4 | 35.8 | 35.9 | -0.80 | 0.333 |  |  |  |  |
|  |
| **S2d) Pitch** |
|  | ***Pitch*** |  |  |  |  |  |  |
|  | ***0.3*** | ***0.5*** | ***0.8*** | ***1.2*** | ***rs*** | ***p*** |  |  |  |  |
| ***Mean Error*** | 13.9 | 14.8 | 10.0 | 13.8 | -0.60 | 0.417 |  |  |  |  |
| ***Image Noise*** | 35.2 | 35.3 | 38.2 | 38.2 | 1.0 | 0.083 |  |  |  |  |
|  |
| **S2e) Reconstruction Kernel** |
|  | ***Kernel*** |  |  |
|  | ***Qr32*** | ***Qr36*** | ***Qr40*** | ***Qr44*** | ***Qr49*** | ***Qr54*** | ***Qr59*** | ***Qr69*** | ***rs*** | ***p*** |
| ***Mean Error*** | 15.4 | 15.4 | 15.6 | 15.8 | 15.7 | 15.4 | 20.4 | 21.5 | 0.83 | 0.015 |
| ***Mean Error (Kernel Qr32 – Qr54 only - posthoc)*** | 15.4 | 15.4 | 15.6 | 15.8 | 15.7 | 15.4 | - | - | 0.6 | 0.242 |
| ***Image Noise*** | 24.4 | 29.3 | 35.5 | 43.6 | 52.7 | 67.0 | 82.9 | 125.6 | 1.0 | <0.001 |

Results of measurement of mean error and image noise in all experiments.

a) At 100 kV / Sn150 kV the lowest possible dosage was at a CTDIvol of 1.5 mGy. b) At 80 kV / 140 kV the lowest possible dosage was at a CTDIvol of 3 mGy; *CTDIvol* volume computed tomography dose index; *rs* Spearman rank correlation coefficient; *Sn* Indicates use of 0.6mm tin filter;